Supplemental Information

Table A1 – Internal standard (ISTD) and derivatization standard (DSTD) used for the analysis of Diesel-CAST samples.

| ISTD | DSTD |
| --- | --- |
| Acenaphthylene-D8 | Fumaric Acid-D2 |
| Acenaphthene-D10 | Adipic Acid-D10 |
| Biphenyl-D10 | P-Nitrophenol-D4 |
| Fluorene-D10 | Glycerol-D8 |
| Anthracene-D10 | Levoglucosane 13C6 |
| Phenanthrene-D10 | Vanillin 13C6 |
| Fluoranthene-D10 | Palmitic Acid-D31 |
| Pyrene-D10 | Dodecanol-D25 |
| Benz[a]anthracene-D12 | Cholesterol-D6 |
| Chrysene-D12 |  |
| Benz[b]fluoranthene-D12 |  |
| Benz[k]fluoranthene-D12 |  |
| Benz[e]pyrene-D12 |  |
| Benz[a]pyrene-D12 |  |
| Perylene-D12 |  |
| Indeno[1,2,3-cd]pyrene-D12 |  |
| Dibenz[a,h]anthracene-D14 |  |
| Benz[ghi]perylene-D12 |  |
| Coronene-D12 |  |
| n-Hexadecane-D34 |  |
| n-Docosane-D46 |  |
| n-Tetracosane-D50 |  |
| n-Triacontane-D62 |  |
| Benz[a]anthracene-7,10-dione-D10 |  |
| 9,10-Anthraquinone 13C6 |  |

Table A2 – Flow parameters. Carrier gas used: Helium.

| Column flow [mL min-1] | Split flow [mL min-1] | t [s] |
| --- | --- | --- |
| 0.1 | 10 | 1200 |
| 1 | 100 | 90 |
| 2.6 | 0 | 600 |
| 1 | 100 | 2510 |

Table A3 – Injection temperature profile.

| T rate [°C s-1] | T [°C] | t [s] |
| --- | --- | --- |
| - | 70 | 40 |
| 2 | 300 | 1350 |
| - | 250 | 3010 |

Table A4 – Column setup.

| Column setup | Type | Length [m] |
| --- | --- | --- |
| Pre-column | BPX5 (0.25 mm i.d. x 0.25 µm df) | 2 |
| 1st Dimension | BPX50 (0.25 mm i.d. x 0.25 µm df) | 58 |
| 2nd Dimension | BPX1 (0.1 mm i.d. x 0.1 µm df) | 1.4 |

Table A5 – GC oven temperature profile. The secondary oven had a temperature offset of 5 °C relative to the GC oven. The modulator temperature offset was 15 °C relative to the secondary oven. The modulation time was 3 s.

| T rate [°C min-1] | T [°C] | Holding time [min] |
| --- | --- | --- |
| - | 40 | 5 |
| 5 | 300 | - |
| - | 300 | 15 |

Table A6 – Mass spectral parameters.

| MS Parameters | Settings |
| --- | --- |
| Transferline temperature | 280 °C |
| Acquisition rate | 100 Hz |
| Acquisition range | 29 – 750 m/z |
| Optimized Voltage Offset | 100 |
| Electron ionization | 70 eV |
| Ion source temperature | 250 °C |

Table A7 – Processing parameters of Leco ChromaTOF. Version 4.50.8.0 optimized for Pegasus.

| Baseline Correction |  |  |  |
| --- | --- | --- | --- |
| Baseline Offset | 1 |  | auto no. of data points for smoothing |
| S/N | 50 |  |  |
| Peak Find | 1st dimension |  | 2nd dimension |
| Peak Width [s] | 12 |  | 0.1 |
| Match required to combine Peaks | - |  | 800 |
| Subpeak | min. S/N |  | 5 |
| Library Search Settings |  |  |  |
| Identity/Search Mode |  | normal/forward |  |
| No. of Hits |  | 10 |  |
| Molecular Weight Min./Max. |  | 29/800 |  |
| Mass Threshold |  | 50 |  |
| Similarity Match Min. |  | 650 |  |
| Library |  | NIST |  |
| Mass for Area Calculation |  | DTIC |  |
| Export |  | .csv-files |  |