# Supplementary Information

# RPLC- and HILIC-based non-targeted metabolomics workflow for blood microsamples

Pauline Couacault1 , Michael Witting1,2,*

1 Metabolomics and Proteomics Core, Helmholtz Zentrum München, Ingostädter Landstraße 1, 85764 Neuherberg, Germany

2 Chair of Analytical Food Chemistry, TUM School of Life Sciences, Technical University of Munich, Maximus-von-Imhof-Forum 2, 85354 Freising-Weihenstephan, Germany

* corresponding author: Dr. Michael Witting, michael.witting@helmholtz-munich.de

ORCiD

Pauline Couacault: 0009-0008-3044-4135

Michael Witting: 0000-0002-1462-4426

#### SI Table S1: Standards used for HILIC method development and optimization

| Standard | Abbreviations | Formula | Comments | Manufacturer |
| --- | --- | --- | --- | --- |
| Alanine | Ala | C3H7NO2 |  | Sigma-Aldrich (Sigma-Aldrich Chemie GmbH, Taufkirchen, Germany) |
| Arginine | Arg | C6H14N4O2 |  |  |
| Asparagine | Asn | C4H8N2O3 |  |  |
| Aspartic acid | Asp | C4H7NO4 |  |  |
| Cadaverine |  | C5H14N2 |  |  |
| Citrulline |  | C6H13N3O3 |  |  |
| Cysteine | Cys | C3H7NO2S |  |  |
| Glutamic acid | Gln | C5H9NO4 |  |  |
| Glutamine | Glu | C5H10N2O3 |  |  |
| Glycine | Gly | C2H5NO2 |  |  |
| Histidine | His | C6H9N3O2 |  |  |
| Isoleucine | Ile | C6H13NO2 |  |  |
| Leucine | Leu | C6H13NO2 |  |  |
| Lysine | Lys | C6H14N2O2 |  |  |
| Methionine | Met | C5H11NO2S |  |  |
| Ornithine |  | C5H12N2O2 |  |  |
| Phenylalanine | Phe | C9H11NO2 |  |  |
| Proline | Pro | C5H9NO2 |  |  |
| Putrescine |  | C4H12N2 |  |  |
| Serine | Ser | C3H7NO3 |  |  |
| Spermidine |  | C7H19N3 |  |  |
| Spermine |  | C10H26N4 |  |  |
| Threonine | Thr | C4H9NO3 |  |  |
| Tryptophan | Trp | C11H12N2O2 |  |  |
| Tyrosine | Typ | C9H11NO3 |  |  |
| Valine | Val | C5H11NO2 |  |  |
| Acetyl-CoA |  | C23H38N7O17P3S | TCA Cycle Metabolite Library |  |
| Oxaloacetic acid |  | C4H4O5 |  |  |
| Malic acid |  | C4H6O5 |  |  |
| CoA |  | C21H36N7O16P3S |  |  |
| Propionyl-CoA |  | C24H40N7O17P3S |  |  |
| Butyryl-CoA |  | C25H42N7O17P3S |  |  |
| Lactic acid |  | C3H6O3 |  |  |
| cis-Aconitic acid |  | C6H6O6 |  |  |
| Adenosine monophosphate | AMP | C10H14N5O7P |  |  |
| Cyclic adenosine monophosphate | cAMP | C10H12N5O6P |  |  |
| Fructose-1,6-bis-Phosphate | F16P | C6H14O12P2 |  |  |
| Ribose-5-Phosphate | R5P | C5H11O8P |  |  |
| Glutathione | GSH | C10H17N3O6S |  |  |
| Glutathione disulfide | GSSG | C20H32N6O12S2 |  |  |
| Riboflavine |  | C17H20N4O6 |  |  |
| Fumaric acid |  | C4H4O4 | Organic acids kit | Supelco (Sigma-Aldrich Chemie GmbH, Taufkirchen, Germany) |
| Succinic acid |  | C4H6O4 |  |  |
| Citric acid |  | C6H8O7 |  |  |
| Isocitric acid |  | C6H8O7 |  |  |
| L-Carnitine | C0 | C7H15NO3 | Acylcarnitines Mix 1 |  |
| O-Acetyl-L-carnitine | C2 | C9H17NO4 |  |  |
| Butyryl-L-carnitine | C4 | C11H21NO4 | Acylcarnitines Mix 2 |  |
| Isobutyryl-L-carnitine | iC4 | C11H21NO4 |  |  |
| Valeryl-L-carnitine | C5 | C12H23NO4 |  |  |
| Isovaleryl-L-carnitine | iC5 | C12H23NO4 |  |  |
| Hexanoyl-L-carnitine | C6 | C13H25NO4 |  |  |
| Octanoyl-L-carnitine | C8 | C15H29NO4 |  |  |
| Decanoyl-L-Carnitine | C10 | C17H33NO4 |  |  |
| Lauroyl-L-carnitine | C12 | C19H37NO4 |  |  |
| Myristoyl-L-carnitine | C14 | C21H41NO4 |  |  |
| Stearoyl-L-carnitine | C18 | C25H49NO4 |  |  |
| trans-2-octenoyl-L-carnitine | C8:1 | C15H27NO4 | Acylcarnitines Mix 3 |  |
| trans-2-decenoyl-L-carnitine | C10:1 | C17H31NO4 |  |  |
| trans-2-dodecenoyl-L-carnitine | C12:1 | C19H35NO4 |  |  |
| trans-2-tetradecenoyl-L-carnitine | C14:1 | C21H39NO4 |  |  |
| trans-2-hexadecenoyl-L-carnitine | C16:1 | C23H43NO4 |  |  |
| oleoyl-L-carnitine | C18:1 | C25H47NO4 |  |  |
| cis-cis-5,8-tetradecandienoyl-L-carnitine | C14:2 | C21H37NO4 |  |  |
| cis,cis-9,12-octadecadienoyl-L-carnitine | C18:2 | C25H45NO4 |  |  |
| Isobutyryl-CoA |  | C25H42N7O17P3S |  | Cayman Chemical (Biomol GmbH, Hamburg, Germany) |
| Fructose-6-Phosphate | F6P | C6H13O9P |  |  |
| Glucose-6-Phosphate | G6P | C6H13O9P |  |  |
| S-Adenosylhomocysteine | SAH | C14H20N6O5S |  |  |
| S-Adenosylmethionine | SAM | C15H23N6O5S |  |  |
| Pyruvic acid |  | C3H4O3 |  | Santa Cruz Biotechnology (Santa Cruz Biotechnology Inc., Heidelberg, Germany) |
| Adenosine diphosphate | ADP | C10H15N5O10P2 |  |  |
| Adenosine triphosphate | ATP | C10H16N5O13P3 |  | Carl Roth (Carl Roth GmbH, Karlsruhe, Germany) |

#### SI Table S2: MS parameters

| Parameter | RPLC-MS positive | RPLC-MS negative | HILIC-MS positive | HILIC-MS negative |
| --- | --- | --- | --- | --- |
| Method duration | 12.1 min | 12.1 min | 10 min | 10 min |
| Source and Gas Parameters |  |  |  |  |
| Ion source gas 1 | 45 psi | 45 psi | 45 psi | 45 psi |
| Ion source gas 2 | 45 psi | 45 psi | 45 psi | 45 psi |
| Curtain gas | 40 psi | 40 psi | 40 psi | 40 psi |
| CAD gas | 7 | 7 | 7 | 7 |
| Temperature | 500°C | 500°C | 500°C | 500°C |
| Experiment |  |  |  |  |
| Experiment | IDA | IDA | IDA | IDA |
| Polarity | Positive | Negative | Positive | Negative |
| Spray voltage | 5500 V | -4500 V | 5500 V | -4500 V |
| TOF MS |  |  |  |  |
| TOF start mass | 70 Da | 70 Da | 70 Da | 70 Da |
| TOF stop mass | 1500 Da | 1500 Da | 1500 Da | 1500 Da |
| Accumulation time | 0.1 s | 0.1 s | 0.1 s | 0.1 s |
| Declustering potential | 80 V | -80 V | 80 V | -80 V |
| DP spread | 0 V | 0 V | 0 V | 0 V |
| Collision energy | 10 V | -10 V | 10 V | -10 V |
| CE Spread | 0 V | 0 V | 0 V | 0 V |
| TOF MS - Advanced Experiment Settings |  |  |  |  |
| Time bins to sum | 6 | 6 | 6 | 6 |
| Channel 1 | TRUE | TRUE | TRUE | TRUE |
| Channel 2 | TRUE | TRUE | TRUE | TRUE |
| Channel 3 | TRUE | TRUE | TRUE | TRUE |
| Channel 4 | TRUE | TRUE | TRUE | TRUE |
| Override QJet RF value | FALSE | FALSE | FALSE | FALSE |
| IDA Criteria | Small molecule | Small molecule | Small molecule | Small molecule |
| Maximum Candidate ions | 12 | 12 | 12 | 12 |
| Intensity treshold exceeds | 100 cps | 100 cps | 100 cps | 100 cps |
| Dynamic background subtraction | TRUE | TRUE | TURE | TRUE |
| Exclude former candidate ions | TRUE | TRUE | TRUE | TRUE |
| For | 1 s | 1 s | 1 s | 1 s |
| After | 1 occurrences | 1 occurrences | 1 occurrences | 1 occurrences |
| Exclusion List | TRUE | TRUE | TRUE | TRUE |
| Exclude isotopes +/- | 4 Da | 4 Da | 4 Da | 4 Da |
| Mass tolerance +/- | 50 mDa | 50 mDa | 50 mDa | 50 mDa |
| TOF MSMS |  |  |  |  |
| Fragmentation mode | CID | CID | CID | CID |
| TOF start mass | 50 Da | 50 Da | 50 Da | 50 Da |
| TOF stop mass | 1500 Da | 1500 Da | 1500 Da | 1500 Da |
| Accumulation time | 0.025 s | 0.025 s | 0.025 s | 0.025 s |
| Declustering potential | 80 V | -80 V | 80 V | -80 V |
| DP Spread | 0 V | 0 V | 0 V | 0 V |
| Zeno pulsing | TRUE | TRUE | TRUE | TRUE |
| Q1 resolution | Unit | Unit | Unit | Unit |
| Collision energy | 35 V | -35 V | 35 V | -35 V |
| CE spread | 15 V | 15 V | 15 V | 15 V |
| TOF MSMS - Advanced Experiment Settings |  |  |  |  |
| Time bins to sum | 6 | 6 | 6 | 6 |
| Channel 1 | TRUE | TRUE | TRUE | TRUE |
| Channel 2 | TRUE | TRUE | TRUE | TRUE |
| Channel 3 | TRUE | TRUE | TRUE | TRUE |
| Channel 4 | TRUE | TRUE | TRUE | TRUE |
| Zeno treshold | 80000 cps | 80000 cps | 80000 cps | 80000 cps |
| Override QJet RF value | FALSE | FALSE | FALSE | FALSE |

#### SI Table S3: RPLC parameters (Sciex ExionLC AD)

| Parameter | Value |
| --- | --- |
| AD Pump |  |
| Stop time | 15.00 min |
| Flow | 0.500 mL/min |
| A.Conc | 95.0% |
| B.Conc | 5.0% |
| B.Curve | 0 |
| Pressure limits Minimum | 20 bar |
| Pressure limits Maximum | 1100 bar |
| Gradient | Advanced |
|  | 1.50 min 0.500 mL/min 95.0%A 5.0%B |
|  | 10.00 min 0.500 mL/min 0.1%A 99.9%B |
|  | 12.00 min 0.500 mL/min 0.1%A 99.9%B |
|  | 12.10 min 0.500 mL/min 95.0%A 5.0%B |
| Solenoid valve |  |
| Pump A: Solvent Valve | A-A-A |
| Compressibility settings |  |
| Compressibility settings | TRUE |
| Mobile Phase A | Water 0.45/GPa |
| Mobile Phase B | Acetonitrile 1.20/GPa |
| AD Autosampler |  |
| Sampling speed | 5.0 µL/s |
| Cooler temperature | TRUE, 5°C |
| AD Autosampler - Rinse settings |  |
| Rinse type | Internal & external |
| Rinse mode | Before and after aspiration, Dip time:0s |
| Rinse pumd method | Rinse port only, Time: 2s |
| Rinse port liquid | R1 |
| Rinsing speed | 35 µL/s |
| Rinse volme | 500 µL |
| Measuring line purge volume | 600 µL |
| Rinsing start time | Specify start time 12.1 min |
| Rinse sequence | R2 --> R0 |

#### SI Table S4: HILIC parameters (Agilent 1290 Infinity II Bio)

| Parameter | Value (Positive) | Value (Negative) |
| --- | --- | --- |
| Binary Pump |  |  |
| Stop time | 15.00 min | 15.00 min |
| Flow | 0.500 mL/min | 0.500 mL/min |
| A.Conc | 0.0% | 0.0% |
| B.Conc | 100.0% | 100.0% |
| B.Curve | 0 | 0 |
| Pressure limits Minimum | 20 bar | 20 bar |
| Pressure limits Maximum | 1300 bar | 1300 bar |
| Gradient | Advanced | Advanced |
|  | 2.00 min 0.500 mL/min 0%A 100%B | 2.00 min 0.500 mL/min 0%A 100%B |
|  | 7.50 min 0.500 mL/min 90%A 10%B | 7.50 min 0.500 mL/min 90%A 10%B |
|  | 9.00 min 0.500 mL/min 90%A 10%B | 9.00 min 0.500 mL/min 90%A 10%B |
|  | 10 min 0.500 mL/min 0%A 100%B | 10 min 0.500 mL/min 0%A 100%B |
| Use Solvent Types | TRUE | TRUE |
| Mobile Phase A | Water + 10mM AF + 0.1% FA | Water + 10mM AA, pH9 |
| Mobile Phase B | Acetonitrile/Water (90:10) + 10mM Ammo Formate + 0.1% Formic Acid | Acetonitrile/Water (90:10) + 10mM Ammo Acetate, pH9 |
| Multisampler |  |  |
| Draw speed | 100 µL/min | 100 µL/min |
| Eject speed | 100 µL/min | 100 µL/min |
| Wait time after drawing | 1.2 s | 1.2 s |
| Injection Volume | 5 µL | 5 µL |
| Cooler temperature | TRUE, 5°C | TRUE, 5°C |
| Rinse settings |  |  |
| Needle Wash Mode | Flush Port | Flush Port |
| Mode | Standard Wash | Standard Wash |
| Duration | 3 s | 3 s |
| Column Comp. |  |  |
| Oven temperature | 40°C | 40°C |